# Shimadzu Philippines Corporation

# Medical System Division

CALIBRATION REPORT

| ## 1. FACILITY INFORMATION |  |  |
| --- | --- | --- |
| ## Facility Name: |  |  |
| ## Address: |  |  |
| ## Telephone/Mobile No.: |  |  |
| ## Email Address: |  |  |
| ## Location within the Facility: |  |  |
| ## 2. MACHINE DETAILS |  |  |
| ## 2.1 Manufacturer: |  |  |
| ## 2.2 Modality: |  |  |
| ## 2.3 Model: |  |  |
| ## 2.4 Serial Number: |  |  |
| ## 2.5 Control Console: | Model: | Serial Number: |
| ## 2.6 X-ray Tube/s Assembly: | Model (1): | Serial Number (1): |
|  | Model (2): | Serial Number (2): |
| ## 2.7 Date of Installation: |  |  |
| ## 3. TECHNICAL SPECIFICATIONS |  |  |
| ## 3.1 Maximum Tube Current mA: |  |  |
| ## 3.2 Maximum Tube Voltage kV: |  |  |
| ## 3.3 Tube Current X time mAs range: |  |  |
| ## 3.4 Tube Voltage kVp range: |  |  |
| ## 3.5 Exposure Time Setting range: |  |  |
| ## 3.6 Maximum Rated Power kW: |  |  |
| ## 3.7 Power Supply: |  |  |
| ## 3.8 Total Inherent Filtration: |  |  |

| ## 4. MECHANICAL CHECKS |  |
| --- | --- |
| ## 4.1 Collimation Assessment | - The total misalignment shall be within +-2% of SID. Each shall be within +-1% of SID - The misalignment between X-Ray field and image receptor shall be within 1.5 degrees - The X-Ray beam perpendicularity shall be within 1.5 degrees. RESULT:_______________________________ |
| ## 5. GENERATOR CHECKS |  |
| ## 5.1 Overload Circuit Protection | Present? Functioning? RESULT:____________________________ |
| ## 5.2 kVp Accuracy | - For 100 kVp and below: The actual kVp shall be within +-6% of the set kVp - For 101 kVp and above: The actual kVp shall be within +-6% of the set kVp RESULT: _____________________________________________________ |
| ## 5.3 mA Accuracy /mA Linearity | - The overall coefficient of linearity shall be <-0.1 RESULT: _____________________________________________________ |
| ## 5.4 Exposure Time Accuracy | - The actual time shall be within +-10% of the set time. For exposure time less than 100ms, +-20% RESULT: _____________________________________________________ |
| ## 6. CALIBRATION TEST DETAILS |  |
| ## 6.1 Date of Machine Calibration: |  |
| ## 6.2 Next Calibration Date: |  |
| ## 6.3 Test Tool Manufacturer: |  |
| ## 6.4 Test Tool Model: |  |
| ## 6.5 Test Tool Serial Number: |  |
| ## 6.6 Test Tool Calibration Date: |  |
| ## 6.7 Name and Signature of Service Engineer: |  |

# Shimadzu Philippines Corporation

# Medical System Division

AVERAGE EXPOSURE OUTPUT

Reference Settings

| FOCAL SPOT: SMALL |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| FOCAL SIZE: 0.6 |  |  |  |  |  |  |
| SID: 100cm |  |  |  |  |  |  |
| Nominal kVP Settings | Measured kVP | mA / mAs | Dose (uGy) | Dose Rate (mGy/s) | Time Settings (msec) | Measured Exposure Time (msec) |

| FOCAL SPOT: LARGE |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| FOCAL SIZE: 1.2 |  |  |  |  |  |  |
| SID: 100cm |  |  |  |  |  |  |
| Nominal kVP Settings | Measured kVP | mA / mAs | Dose (uGy) | Dose Rate (mGy/s) | Time Settings (msec) | Measured Exposure Time (msec) |

| PERFORMANCE CRITERIA | TEST RESULT |
| --- | --- |
| The average kVp shall not differ from the nominal kVp by ±6% for voltages less than or equal to 100kVp, or 6kV for voltages greater than 100kVp. |  |
| The actual time shall be within +-10% of the set time. For exposure time less than 100ms, +-20% |  |